## Complaince
TECHNICAL COMPLIANCE FORM
Ref: Tender No:
TATA MEMORIAL CENTRE
ADVANCED CENTRE FOR TREATMENT, RESEARCH & EDUCATION IN CANCER
TECHNICAL SPECIFICATION OF :- - " Anesthesia Delivery System -1 Number "
Intended Use: For giving anesthesia to the patient, those who are going to undergo the surgical procedure in CT scan
Sr. No. | Name of the Vendor :
1 | Technical Specifications / Scope of supply | Quantity | COMPLIANCE 1) Mention "Complied /Not Complied" 2) Highlight any deviations. 3) Mention part number/Catalogue number of the relevant item quoted. | COMPLIANCE 1) Mention "Complied /Not Complied" 2) Highlight any deviations. 3) Mention part number/Catalogue number of the relevant item quoted. | COMPLIANCE 1) Mention "Complied /Not Complied" 2) Highlight any deviations. 3) Mention part number/Catalogue number of the relevant item quoted.
2 | Name of the Manufacturer :
3 | Model Name / Model No. of the Equipment (If the equipment is a combination of more than one equipment then list down the model no. pertaining to each) Can Quote more than one model.:
4 | Year of Introduction Internationally
5 | Year of Introduction in India
6 | Technical Specifications/Scope of supply
7 | " Anesthesia Delivery system" , complete as per below mentioned configuration and specifications: | Qty: 1 No.
8 | Features and specifications required:
9 | Durable with a sturdy kit.
10 | Antistatic wheels with brakes
11 | Shelves for monitors
12 | The anesthesia workstation should have in built ventilator with min. 6'' TFT display.
13 | Gas delivery system:
14 | Gas specific high pressure blocks.
15 | Non interchangable gas supply inlet for cylinders with step down regulators
16 | Pressure reducing valves.
17 | Color coded indexed gas supply hoses.
18 | Provision of additional O2 port for powering a ventury device.
19 | Provision for manual as well as mechanical ventilation.
20 | Provision for both open as well as closed circuit.
21 | Provision for attachement of 2 cylinders one each of O2 and N2O cylinders.
22 | Provision to monitor the expired gas volume, O2 monitor sensor probe along with all the tubings and connections with alarms.
23 | Ability to monitor the expired TV/minute volume,respiratoy rate and FiO2
24 | Quick change over form open circuit to closed cicuit and vice versa.
25 | Ability to bypass the soda lime canister and convert to an open circuit.
26 | Provision for emergency flow of 100% O2 under high pressure through both circuits.
27 | Audible and visual alarms for airway pressure, disconection, oxygen failure.
28 | Flow meters:
29 | For oxygen, nitrous oxide and air with safety feature
30 | Oxygen failure alarm, audible and visible
31 | Anti-hypoxia safety mechanism to ensure that O2 concentration doesn’t fall below minimum 23% at any time.
32 | Vaporisers:2Nos with each Unit
33 | Should have provision to connect two vaporisers.
34 | Vaporiser safety lock.
35 | Vendor Should provide Vaporisers-2Nos (isoflurane-1No,sevoflurane-1No)
36 | Ventilators:
37 | Ventilators for controlled ventilation of adults as well as children
38 | Volume contorlled/pressure controlled/Pressure Support.
39 | Inbuilt PEEP administration.
40 | Provision for manual and mechanical ventilation
41 | Monitoring for airway prsesure- Peak and PEEP
42 | Provision to monitor expired gas volume, O2 monitor sensor probe along with all the tubings and connections with alarms.
43 | Ability to monitor the expired TV/minute volume, tidal volume, respiratoy rate and FiO2
44 | Tidal volume delivery should be in the range from 30ml(or less) to 1400 ml (or more)
45 | Respiration rate: 4-60 bpm (or more)
46 | PEEP setting: upto 20 cmH20
47 | Peak pressure setting upto 50cmH20
48 | Breathing system:
49 | Should have integrated CO2 absorber with easy removal and refitting during the operation.
50 | Should be autoclavable at 134 deg C.
51 | Should have APL valve, inspiratory valve, expiratory valve.
52 | should have mechanism to prevent moisture condensation (eg. Heating circuit or equivalent )
53 | Should work on mains Power and with battery backup of 45 minutes minimum with fully charged batteries
54 | International standards : The unit should comply with international standards and should have CE or FDA marking.
55 | Standard accessories (Which comes along with the machine, pricing of which is part of the equipment)
56 | Should provide all standard acceossories
57 | Open and closed circuit for both adult and pediatric
58 | Essential accessories (List down essential accessories to carry out certain functions but comes at an extra cost. Prices to be quoted separately)
59 | List down all optional accessroies with price in price bid
60 | Essential Spares: Please list down the spare parts which requires to changed during Preventive/ Break down maintenance:
61 | Essential consumables:
62 | Indicate if the quoted model needs proprietary consumables (Equipment being closed or open system)
63 | Provide List of consumables with their prices in the Financial bid. (partial list is given below but all relevant consumable prices are to be quoted)
64 | Upgradibility capability (List down possible upgrades for the quoted model)
65 | On-Line UPS/Voltage Stabilizer requirements
66 | Periodic Calibration
67 | Regulatory Approvals, If any. Details with copies of approvals.
68 | Safety requirements:It should follow International / national safety requirement. Please specify certification with certifying agency and country with copies of certificate.
69 | User's list: A list of installations with the address and contact numbers to be provided. ( User list should be for the quoted model.)
70 | Multiple Models: If more than one model is offered, then they are to be separately filled and submitted according to the given format.
71 | Input power supply requirements: 240 V AC ± 10 %, 50 Hz, single phase. Specify if any other power supply requirement is recommended.
72 | After Sales Service Support.
73 | 1. Complete operational manuals and technical information including circuit diagrams should be provided. 2. Company should provide technical support for hardware and software, spares and consumables. 3. Company should provide required spares and consumables for at least 7 years even after discontinuation of the model. 4. Bidders who do not qualify this critera shall not be considered. Bidder should provide address of the authorized service centre and details of number of personnel with training credentials.
74 | Preinstallation Requirements :- Provide details of preinstallation requirements, special ambient conditions, if any, space requirement, any other special needs.
75 | Installation, Commissioning, testing and Training
76 | Unpacking and Shifting the consignment to the installation site is to be included in the scope of supply. Bidder/manufacturer/authorized service provider should take responsibility to lift/shift the consignment from unloading site to the installation site . Unloading site shall be "Stores Department, KS Building, ACTREC Campus". If needed, Bidder has to arrange for the labourers at no charge to ACTREC. ( Before submitting the quotation, bidders may visit ACTREC to know unloading site and installation site)
77 | Installation,Commissioning and Training is included in the scope of supply. Bidder, Manufacturer and/or its authorized representative should undertake installation and commissioning of the equipment.
78 | Complete system should be installed, tested for its performance as per manufacturer's SOP/guidelines and demonstrated to the Institute's Users. In depth training should be provided to the Institute's users for maintenance,usage and applications.
79 | Warranty and after sales support
80 | Warranty: 2 years comprehensive-on site- Factory warranty- for entire equipment. All third party major items, accessories/attachment etc. integrated with the system should also cover 2 years -factory warranty from respective manufacturers.
81 | Quote for AMC and CMC costs separately. AMC and CMC cost for 5 years-post warrant,
82 | Uptime guarantee during warranty period and CMC : Bidder should specify the uptime guarantee offered during the warranty and CMC period.
83 | Quote optionally for the following items: ( Price should be separately quoted in the financial bid)
84 | List of spares and consumables recommended to be kept in stock |  | Quoted/Not Quoted
85 | List of all optionally available accessories |  | Quoted/Not Quoted
86 | Important terms to be noted by the bidders:
87 | Read the above scope of supply carefully and quote accordingly. Incomplete and /or partially complete offers are liable to be rejected. |  | Agree/Not Agree
88 | Manufacturer/bidder should clearly indicate typical life span / Life expectancy of the quoted equipment. |  | -- years
89 | Mention typical site requirements for the installation. |  | --
90 | Mention the time required to install the system. |  | - days
91 | Quote separately for all models meeting the described technical specifications including the latest model. |  | Agree/Not Agree
92 | Bidders should clearly mention about compliance of the warranty and AMC/free service terms in the technical bid. Bids not fully complying to the warranty and AMC/free service terms are liable to be rejected. |  | Agree/Not Agree
93 | Item/Equipment quoted should not be a second hand/refurbished/on verge of obsolescence/outdated model. Similarly, after award of contract, the item supplied should be a newly manufactured item and not the old item already manufactured and stocked for more than 1 month. |  | Agree/Not Agree
94 | After opening of the Technical bid (Part-1), Physical demonstration of the quoted model may have to be shown / arranged by the bidder, if requested by the Institute. Physical demonstration may be shown at one of the end user’s site/Principle company’s application lab/manufacturing site located in Mumbai/Navi Mumbai/Thane cities. If there are no installations of the quoted model in Mumbai/Navi Mumbai/Thane cities, then the quoted model may have to be brought in at ACTREC for demo purpose within 10 days from the date of request. Physical Demonstration may be requested to confirm the availability of any or all technical features as mentioned/stated in the technical bid. Physical Demonstration will also be a part of technical evaluation process. If the bidder does not comply, such bids are liable to be disqualified. (Demonstration of quoted model is to be shown and not the demonstration of similar models with different technical specifications and features) |  | Agree/Not Agree
95 | Past experience of the bidders in terms of quality of supplied equipments, after sales service and application support will be taken into consideration while technical evaluation.. Bidders who has unsatisfactory past experience in last 2-3 years, in terms of quality of supplied equipments, after sales service and application support, bids of such bidders may liable to be rejected. |  | Agree/Not Agree
96 | Complete and detailed information should be provided in respect to each point specified in the specifications. Technical bids that are incomplete in any respect are liable to be rejected. Provide relevant supportive information, publications, catalogue, etc. Bidders providing misleading or wrong information are liable to be rejected. All technical claims should be printed in the technical brochure of the equipment. |  | Agree/Not Agree
97 | If any contradictory statements /figures/information is observed in the compliance chart and in the technical bid, then the technical information mentioned in the product literature/brochure will be considered true and further evaluation will be done based on the information given in the product literature/brochure. Regarding the warranty and free AMC terms, information quoted in the technical bid will only be taken as valid terms. |  | Agree/Not Agree
98 | Compliance column may be filled with relevant data, figures, range etc. as applicable. Do not just mention "YES / NO / Complied ". |  | Agree/Not Agree
99 | Decleration by the bidder
100 | We have quoted for all the items meeting the description/scope of supply in the Financial bid as per prescribed format of the Tender documents and we agree that Partial/incomplete offers are liable for rejection. |  | Yes/No

## vendorcap
ANNEXURE -G
Ref: Tender No: ACTREC/IMP/2017-18/E-Tender-24
TATA MEMORIAL CENTRE
ADVANCED CENTRE FOR TREATMENT, RESEARCH & EDUCATION IN CANCER (ACTREC)
PURCHASE SECTION
Plot No.1&2, Sector No.22, Kharghar, Navi Mumbai – 410210
TECHNICAL SPECIFICATION OF "Anesthesia Delivery System"
VENDOR CAPABILITY FORM
Ref: Tender No:
1 | Name / Title of the Bidder
2 | Due Date of the submission of the tender.
3 | Receipt No. for EMD Amount and Date
4 | Name of the Item / Work
5 | Full Address (recent)
E-Mail (recent)
Tel. No & Mobile No. (recent)
Fax (recent)
6 | Name of the person authorized to deal / undertake business for and on behalf of the bidder
Tel. No & Mobile No. (Recent)
Fax (Recent)
E-Mail (Recent)
7 | Legal entity of the bidder whether Firm / Society / Company / Other entity
a. Registration No.
b. Authority with whom registered
c. Licence No. granted by for
9 | Authorized Area of operation in India
10 | Name of the Principal Organization / Company for and on behalf working in India
11 | Origin of the Principal Organization / Company
Tel. No. & Mobile No.
Fax
13 | Name & Address of the Bankers of the bidders.
14 | Authority / Delegation / Licence No. & Date granted by the principal to the representative bidder
15 | PAN No.
TAN No.
16 | Registration No. granted by Central Excise Commissioner for Service tax
17 | GST No
18 | Central Excise License No.
19 | Import / Export Code No.
License No. for import
20 | No. of staff employed per equipment
a. Scientific
b. Technical
c. Administrative
d. Finance
21 | Support facility equipment No.
22 | Experience of the bidder in dealing with the tendered item. Tenderer must have similar job done in the line of business / experience with 3 to 5 years will be considered
23 | Whether supply of any item / service to TMC in past; if yes indicate the Purchase Order No. & Date

## NEFT
NEFT FORM
Ref: Tender No:
ANNEXURE-J
1) | Name of the Vendor
2) | Vendor Address & Other Particulars
a) | PAN NO.
b) | GST NO.
d) | Mobile No.
e) | Email ID
3) | Account Holder's Name (Title of the Account)
4) | Bank Account No.
5) | Bank Name, Branch & Address
6) | 9-Digit MICR code of the bank
7) | Account type (SB/CURRENT)
8) | IFSC Code (attach xerox copy of cheque)
I hereby declare that the particulars given above are correct and complet. If the transaction is delayed or not effected at all for reasons of incomplete information, I would not hold the user institution responsible. I agree to discharge the responsibilies as a participant under the scheme.
Signature of the Vendor with seal.
Certified that the particulars furnished above are correct as per our records.
Signature of the authorised official from the bank.
Bank stamp :
Date :
Note : Xerox copy of cheque may be attached,without which the form will not be accepted.